Hi Matt,

Some comments with regard to document Feed Control Commands Ver 3.docx:

1) I will adjust the code to match the pulses-per-rotation of the new fan such that proper RPM is computed. I will also fix the issue Rob pointed out where the RPM does not report 0 when the fan is not rotating at all.

2) gettemp a1 through a6 should be working fine with current firmware. If not, double check the addresses and wiring of the external TC74 devices. Rob should feel free, as he has in the past, to contact me.

3) I will add command getcryoattemp

4) I will add code to return LNA temp. in Kelvin units as specified

5) I will add getaccel to return accelerometer X,Y,Z statistics

6) setrelay on/off should be available now.

7) I will change getvac (gv) to return units mbar instead of Torr

8) Regarding setcoolmode, setcoolstart, setmaxreject, setcoolerstate: I recommend, except for setting the fan to a safe cooling speed, that the control board does not automatically start any peripheral support equipment (e.g., cryo and vacuum). At the ATA, that should be the responsibility of the higher-level control system like done for other antenna equipment (such as motor relay latch). In the Minex lab, cryo and vacuum start should be done by a human operator. But, it is important for the
control board to implement simple protection algorithms, such as shutting off the cooler based on rejection temperature and so forth (as described on page 5).

9) I may offer a different suggestion for monitoring the diaphragm pump. We agreed in the past that we would track the diaphragm pump duty cycle through power levels.

10) Rather than separate commands (e.g., getmaxa1) I will return statistics (i.e. mean, stddev, min, max) for measured values where appropriate.

11) Regarding gettimeon, gettimecool, gettimevac, and related history commands: the control board does not have a real-time clock or appropriate storage to support these. The ATA control system has a database (MetadataServer?) where these values should be polled from the control boards and stored. I recommend adding a separate section to your document that defines the monitor and control requirements for whomever (Todd ?) at SRI will be adding support to the control system for the Antonio feed.

12) I will change the fan control loop as defined on page 5 and to be parameterized. I do not understand symbols T2/T1 but may figure it out on a second reading.

13) I will update the Python data logging program as described on page 6.

Best,
Rob